GEBRAUCHSINFORMATION	0123

##### HISTO TRAY ABC / AB / B27 RESERVOIR

# IN VITRO DIAGNOSTIKA

## Produktbeschreibung

HISTO TRAY ABC / AB / B27 RESERVOIR sind Vorratsgefäße mit Anti-HLA-ABC- / AB- / B27-Seren und Kontrollen für die Gewebetypisierung der HLA-Klasse I Antigene. Worksheets zur Auswertung und Ergebnislisten sind beigefügt.

## Testprinzip

Anti-HLA-Seren reagieren mit korrespondierenden, membrangebundenen Antigenen von humanen Lymphozyten. Durch den Zusatz von Kaninchen-komplement kommt es zu Strukturveränderungen der Zellmembran, so dass ein Indikatorfarbstoff in die Lymphozyten eindringen und diese anfärben kann (positive Reaktion). Findet keine Antigen-Antikörper-Reaktion statt, bleibt die Zellmembran intakt. Die Zellen können den Farbstoff nicht aufnehmen (negative Reaktion).

Zusätzlich benötigte Materialien

Aqua dest., Paraffinöl

Mikrotestplatten mit 60 oder 72 Kavitäten (z.B. Fa. Greiner), Tropfgerät für 1 – 5 µl, Mikroliterspritzen für 1 – 5 µl

Optional: Vakuumeinschweißgerät, Vakuumbeutel

NIH-Test

Zellkulturmedium (z.B. RPMI 1640), Zelltrennmedium (z.B. HISTOPREP), Zentrifugenröhrchen (12 ml), Pasteurpipetten, Zentrifuge, Neubauer-Zählkammer oder Zellcounter

Kaninchenkomplement

Eosin-Lösung (5% wässrig), Formaldehyd-Lösung (37%, pH 7,2),

Deckgläschen, inverses Phasenkontrastmikroskop

Immuno-Beads-Technik

Immuno Beads (s. Herstellerangaben)

Kaninchenkomplement

Acridinorange/Ethidiumbromid, Quenching-Lösung, EDTA (8% wässrig)

Fluoreszenzmikroskop

Testdurchführung – Tropfen der Mikrotestplatten

- Vorratsgefäß mit den Anti-HLA-ABC- / AB- / B27-Seren und Kontrollen auftauen.

- Mikrotestplatten mit Produktnamen, Lotbezeichnung und Haltbarkeitsdatum wasserfest kennzeichen.

- Mikroliterspritze oder Tropfgerät mindestens 3 x mit jeweils frischem Aqua dest. spülen. Wird ein Tropfgerät eingesetzt, ist die Bedienungsanleitung des 
	Herstellers zu beachten.

- In jede Kavität der Mikrotestplatten 5 µl Paraffinöl tropfen.

- Dann aus dem Vorratsgefäß je 1 µl Anti-HLA-Serum bzw. Kontrolle in die entsprechende Kavität tropfen.

Es ist darauf zu achten, dass sich das rotgefärbte Serum am Boden der Kavität unter dem Paraffinöl befindet.

- Bis zur Verwendung die betropften Mikrotestplatten bei  -20°C einfrieren.

Für die Lagerung wird empfohlen, die betropften Platten in Vakuumbeutel unter Vakuum einzuschweißen.

## Testdurchführung - Isolierung von Lymphozyten aus z.B. heparinisiertem Blut

1.	Zur Steigerung der Zellausbeute 4 ml heparinisiertes (50 IE/ml) Blut mit 4 ml Zellkulturmedium, z.B. RPMI 1640 verdünnen.

2.	4 - 5 ml Zelltrennmedium, z.B. HISTOPREP in ein Zentrifugenröhrchen (12 ml) geben.

3.	Ca. 6 ml verdünntes Blut mit einer Pasteurpipette am Innenrand des Röhrchens vorsichtig auf den Gradienten schichten.

4.	15 Minuten bei 1.200 x g und einer Temperatur von 18...22°C (Raumtemperatur) zentrifugieren. Zentrifuge ungebremst auslaufen lassen.

5.	Mit einer Pasteurpipette den Lymphozytenring (Interphase) abheben und in ein neues Zentrifugenröhrchen geben.

6.	Zum Waschen der Lymphozyten mit Zellkulturmedium, z.B. RPMI 1640 auffüllen und 10 Minuten bei 550 x g zentrifugieren; Überstand verwerfen; Sediment resuspendieren und mit Zellkulturmedium, z.B. RPMI 1640 auffüllen.

7.	10 Minuten bei 230 x g zentrifugieren; Überstand verwerfen, Sediment resuspendieren und mit Zellkulturmedium, z.B. RPMI 1640 auffüllen.

8.	Nochmals 10 Minuten bei 110 x g zentrifugieren und den Überstand verwerfen.

- Sediment in Zellkulturmedium, z.B. RPMI 1640 so resuspendieren, daß eine Endkonzentration von 2000 - 3000 Lymphozyten pro µl vorliegt (Neubauer-Zählkammer oder Zellcounter).

## Testdurchführung - NIH-Test

1.	Mikrotestplatten mit den vorgetropften Anti-HLA-Seren auf eine Temperatur von 18...22°C (Raumtemperatur) bringen.

- In jede vorgetropfte Kavität 1 µl Lymphozytensuspension (2.000 - 3.000 Zellen) geben.

Um eine korrekte Antigen-Antikörperreaktion zu gewährleisten, ist darauf zu achten, daß sich Zellen und Antiserum verbinden.

3.	30 Minuten bei einer Temperatur von 18...22°C (Raumtemperatur) inkubieren.

4.	5 - 6 µl Kaninchenkomplement zugeben.

5.	60 Minuten bei einer Temperatur von 18...22°C (Raumtemperatur) inkubieren.

6.	3 - 4 µl Eosin-Lösung (5% wässrig) (Softtouchmethode) zugeben und 5 - 10 Minuten inkubieren.

7.	Mit 5 - 6 µl Formaldehyd-Lösung (37%; pH 7,2) fixieren (Softtouchmethode) und mindestens 60 Minuten stehenlassen.

8.	Zur besseren Beurteilung kurz vor dem Mikroskopieren ein Deckgläschen auflegen und mit inversem Phasenkontrastmikroskop auswerten.

Testdurchführung - Isolierung von T-Lymphozyten mit Immunobeads-Technik

Die Isolierung der T-Lymphozyten mit Immuno-Beads-Technik und die zur Färbung und Fixation benötigten Reagenzien sind den Gebrauchsinformationen des Herstellers zu entnehmen.

## Testdurchführung - Immuno-Beads-Technik (IMB)

1.	Mikrotestplatten mit den vorgetropften Anti-HLA-Seren auf eine Temperatur von 18...22°C (Raumtemperatur) bringen.

- In jede vorgetropfte Kavität 1 µl IMB-T-Lymphozytensuspension (ca. 1.000 Zellen) geben.

Um eine korrekte Antigen-Antikörperreaktion zu gewährleisten, ist darauf zu achten, daß sich Zellen und Antiserum verbinden.

3.	30 Minuten bei einer Temperatur von 18...22°C (Raumtemperatur) inkubieren.

4.	5 µl Kaninchenkomplement Acridinorange/Ethidiumbromid (AO/EB) zugeben (1.000 µl Kaninchenkomplement + 20 µl AO/EB).

5.	60 Minuten bei einer Temperatur von 18...22°C (Raumtemperatur) im Dunkeln inkubieren.

6.	5 µl EDTA-/Quenching-Lösung zugeben (2.000 µl Quenching-Lösung + 1.000 µl EDTA 8% wässig).

7.	Unter einem Fluoreszenzmikroskop ablesen.

## Bewertung der Reaktionen

Der Anteil der lysierten Lymphozyten an der Gesamtlymphozytenzahl wird als Scorewert angegeben.

% lysierte Zellen		Bewertung

0  –	  19%	=	Score 1	negativ

20  –	  39%	=	Score 2	fraglich negativ

40  –	  59%	=	Score 4	schwach positiv

60  –	  79%	=	Score 6	positiv

80  –	100%	=	Score 8	stark positiv

=	Score 0	nicht auswertbar

## Fehlerquellen

## Ursache falsch negativer oder schwacher Reaktionen

- Erythrozyten können das Ablesen erschweren

- Kontamination mit Thrombozyten

- Zu hohe Lymphozytenzahl

- Zu wenig Anti-HLA-Serum getropft

- Gelbfärbung der Anti-HLA-Seren

- Platten aufgetaut und wieder eingefroren

- Komplement vor Verwendung zu lange bei Raumtemperatur gelagert

- Reste von aufgelöstem Komplement eingefroren und erneut verwendet

- Zu kurze Inkubationszeiten

- Zu niedrige Inkubationstemperatur

## Ursache falsch positiver Reaktionen

- Kreuzreaktionen

- Zu viel Anti-HLA-Serum getropft

- Zu lange Inkubationszeiten

- Zu hohe Inkubationstemperatur

- Vorgeschädigte Lymphozyten (Negativkontrolle ist positiv = „background“)

- Reaktionen wurden nicht abgestoppt

Kaninchenkomplement

Lyophilisiertes Kaninchenkomplement kurz vor Gebrauch mit 1 bzw. 5 ml (siehe Packungsgröße) Aqua dest. rehydrieren. Die Rehydrierung dauert 10 - 15 Minuten.

Aufgelöstes Kaninchenkomplement muß kühl (2...8°C) gelagert und innerhalb von 3 - 4 Stunden verbraucht werden.

Aufgelöstes Kaninchenkomplement nicht einfrieren!

### Leistungsdaten

Diagnostische Sensitivität und Spezifität (R-Wert) sind den Ergebnislisten HISTO TRAY ABC / AB / B27 RESERVOIR zu entnehmen.

Literatur

Bodmer, J. et al., 1997. Tissue Antigens 49:297-321

### Warn- und Entsorgungshinweise

HISTO TRAY ABC / AB / B27 RESERVOIR sind nur für den in vitro diagnostischen Gebrauch geeignet und sollten nur von geschultem, in Histo-kompatibilitätstestung erfahrenem Fachpersonal angewendet werden. Transfusionsrichtlinien und EFI- / DGI-Standards sind zu beachten, insbesondere bei zweifelhaften Typisierungsergebnissen.

Humanes Ausgangsmaterial für die Produktion der Testreagenzien wurde auf HBsAg und Antikörper gegen HIV und HCV getestet. Nur negatives Material wurde für die Produktion verwendet. Trotzdem sollten sämtliche für den Test verwendete Materialien biologischen Ursprungs, wie Blut, Serumproben und Kontrollseren als potentiell infektiös betrachtet werden, da keine Testmethode alle infektiösen Krankheitserreger nachweisen kann. Beim Umgang mit biologischen Materialien werden deshalb angemessene Sicherheitsvorkehrungen empfohlen (nicht mit dem Mund pipettieren; Schutzhandschuhe bei der Testdurchführung tragen; Händedesinfektion nach der Testdurchführung).

Biologische Materialien müssen vor der Entsorgung inaktiviert werden (z.B. durch Autoklavieren). Einwegmaterialien sind nach Gebrauch zu autoklavieren oder zu verbrennen. Verschüttetes potentiell infektiöses Material sollte unverzüglich mit einem saugfähigen Papiertuch entfernt werden und der kontaminierte Bereich mit einem geeigneten Desinfektionsmittel oder 70%igem Ethanol desinfiziert werden. Material, dass für die Entfernung von Verschüttetem benutzt wurde, muss vor der Entsorgung inaktiviert werden (z.B. durch Autoklavieren).

- Die Anti-HLA-Seren enthalten < 0,1% NaN3 als Konservierungsmittel. In der Konzentration von < 0,1% gilt NaN3 nicht mehr als gesundheitsschädlich, trotzdem sollte ein Kontakt mit der Haut und Schleimhäuten vermieden werden. Kupfer und Blei, die in einigen Rohrsystemen eingesetzt werden, können mit Azid explosive Salze bilden. Die im Reagenz enthaltenen Azidmengen sind klein, trotzdem sollte bei der Beseitigung von Azid-haltigem Material mit reichlich Wasser nachgespült werden.

Die Entsorgung aller Proben und Testmaterialien sollte entsprechend der gesetzlichen Richtlinien erfolgen.

Für Quenching-Lösung, Formaldehyd-Lösung und Acridinorange/Ethidiumbromid (AO/EB) sollten die Warn-und Entsorgungshinweise der Hersteller befolgt werden.

Eine gelbe Verfärbung der Anti-HLA-Seren, die auch nach dem Auftauen bestehen bleibt, zeigt eine Änderung im pH-Wert an. Derartige Platten sollten nicht für den Test eingesetzt werden.

Sicherheitsdatenblätter auf Anfrage erhältlich.

HISTO TRAY ABC/ AB/ B27 RESERVOIR nicht nach Ablauf des auf dem Etikett angegebenen Haltbarkeitsdatums benutzen.

Konservierungsmittel:	< 0,1% NaN3

Haltbarkeit:	bis zum aufgedruckten Datum auf den Etiketten

Lagerung:	 -20°C

Packungsgröße:	gemäß Angaben auf dem Etikett

| # Erklärung der Symbole auf den Etiketten |  |
| --- | --- |
| IVD | In vitro Diagnostikum |
| LOT | Lot-Nr. |
| REF | Bestell-Nr. |
|  | Lagertemperatur |
|  | Verwendbar bis |
|  | Gebrauchsinformation beachten |

|  | BAG-BiologischeAnalysensystemGmbH Amtsgerichtsstr. 1 – 5, D-35423 Lich |
| --- | --- |
|  | Auftragsannahme Customer Service Telefon: (+49) 64 04 / 9 25 450 (+49) 64 04 / 9 25 125 # Fax: (+49) 64 04 / 9 25 460 (+49) 64 04 / 9 25 421 # Internet: www.bag-germany.com |

INSTRUCTIONS FOR USE	0123

HISTO TRAY ABC / AB / B27 RESERVOIR

FOR IN VITRO DIAGNOSTIC USE

###### Description of product

###### HISTO TRAY ABC / AB / B27 RESERVOIR are storage trays with anti-HLA ABC / AB / B27 sera and controls for tissue typing of HLA class I antigens. Worksheets for evaluation and the listing of test results are enclosed.

Test principle

Anti-HLA sera react with the correspondent membrane-bound antigens on human lymphocytes. The addition of rabbit complement results in a structural change of the cell membrane which leads to a penetration of an indicator dye. Stained lymphocytes = positve reaction. In case of missing antigen-antibody reaction, the cell membrane is intact. No penetration of indicator dye takes place and the cells remain unstained = negative reaction.

Additional materials required

Aqua dest., paraffin perliquidum

Microtest plates with 60 or 72 wells (e.g. Greiner), automatic dropper for 1 – 5 µl, microliter syringe for 1 – 5 µl

Optional: Vacuum packing machine, vacuum bags

NIH technique

Cell culture medium (e.g. RPMI 1640), cell separation medium (e.g. HISTOPREP), centrifuge tubes (12 ml), pasteur pipettes, centrifuge, Neubauer count chamber or cell counter

Rabbit complement

Eosin solution (5% aqueous), Formaldehyde solution (37%, pH 7,2)

Cover glasses, inverse phase contrast microscope

Immuno Beads method

Immuno Beads (details see manufacturer)

Rabbit complement

Acridinorange/Ethidiumbromide, quenching-solution, EDTA (8% aqueous)

Fluorescence microscope

Test procedure – Dropping of microtest plates

- Thaw storage tray with anti-HLA sera and controls.

- Label microtest plates with product name, lot and expiry date (water proof labelling).

- Flush microliter syringe or automatic dropper at least 3 x with fresh aqua dest. If using an automatic dropper, note the instruction manual of the 
	manufacturer.

- Drop 5 µl paraffin perliquidum in each well of the microtest plate.

- Drop 1 µl anti-HLA serum or control from the storage tray into the respective well.

Pay attention that the red coloured serum is located under the paraffin perliquidum at the bottom of the well.

- Store the dropped microtest plates at  -20°C up to use.

It is recommended to weld the dropped plates in vacuum bags under vacuum for storage.

## Test procedure - Isolation of lymphocytes from e.g. heparinized blood

1.	In order to increase the cell yield, dilute 4 ml of heparinized blood (50 I.U./ml) with 4 ml of cellculture medium (e.g. RPMI 1640).

2.	Pipet 4 - 5 ml of cell separation medium, e.g. HISTOPREP into a centrifuge tube (12 ml).

3.	Carefully add app. 6 ml of diluted blood with a Pasteur pipette to the gradient alongside the inner edge of the tube.

4.	Centrifuge 15 minutes at 1200 x g and a temperature of 18...22°C (room temperature). Centrifuge without braking.

5.	Take off the lymphocyte ring (interphase) using a Pasteur pipette and pipet it into a new centrifuge tube.

6.	For lymphocyte washing, fill it up with cellculture medium, e.g. RPMI 1640 and centrifuge for 10 minutes at 550 x g. Discard the supernatant, resuspend the sediment and fill it up with cellculture medium, e.g. RPMI 1640.

7.	Centrifuge for 10 minutes at 230 x g, discard the supernatant, resuspend the bottom sediment and fill it up with cellculture medium, e.g. RPMI 1640.

8.	Centrifuge for 10 minutes at 110 x g and discard the supernatant.

- Resuspend the sediment in cellculture medium, e.g. RPMI 1640, and adjust to a final concentration of 2000 - 3000 lymphocytes per µl (Neubauer count chamber or cell counter).

## Testprocedure – NIH technique

1.	Bring the microtest plates with predropped anti-HLA sera to a temperature of 18...22°C (room temperature).

- Place 1 µl lymphocyte suspension (2.000 - 3.000 cells) into each predropped well.

In order to guarantee sufficient antigen-antibody reaction it is necessary that antiserum and cells touch each other.

3.	Incubate at a temperature of 18...22°C (room temperature) for 30 minutes.

4.	Add 5 - 6 µl rabbit complement.

5.	Incubate at a temperature of 18...22°C (room temperature) for 60 minutes.

6.	Add 3 - 4 µl Eosin solution (5% aqueous) (soft touch method) and incubate for 5 - 10 minutes.

7.	Fix with 5 - 6 µl Formaldehyde solution (37%, pH 7.2) (soft touch method). Allow sedimentation of cells  at least 60 minutes.

8.	Cover the tray with a cover glass shortly before reading under an inverse phase contrast microscope.

Test procedure - Isolation of T-lymphocytes from e.g. heparinized blood

Isolation of the T-lymphocytes using the Immuno Beads method as well as staining and fixation reagents according to manufacturers instructions.

Test procedure - Immuno Beads (IMB) method

1.	Bring the microtest plates with predropped anti-HLA sera to a temperature of 18...22°C (room temperature).

- Place 1 µl IMB-T-lymphocyte suspension (app.1.000 cells) into each predropped well.

In order to guarantee sufficient antigen-antibody reaction it is necessary that antiserum and cells touch each other.

3.	Incubate at a temperature of 18...22°C (room temperature) for 30 minutes.

4.	Add 5 µl rabbit complement Acridinorange/Ethidiumbromide (AO/EB) (1.000 µl rabbit complement + 20 µl AO/EB).

5.	Incubate for 60 minutes at a temperature of 18...22°C (room temperature) in darkness.

6.	Add 5 µl EDTA-/quenching-solution (2.000 µl quenching solution + 1.000 µl EDTA 8% aqueous).

7.	Read under a fluorescence microscope.

## Evaluation of results

The amount of lysed lymphocytes compared with the total amount of lymphocytes is quoted as a score value in each well.

% lysed cells	Evaluation

0	–	  19%	=	Score 1	negative

20	–	  39%	=	Score 2	doubtful negative

40	–	  59%	=	Score 4	weak positive

60	–	  79%	=	Score 6	positive

80	–	100%	=	Score 8	strong positive

=	Score 0	no evaluation possible

Troubleshooting

Causes of false negative or weak reactions

- Erythrocyte contamination can make microscopic evaluation difficult

- Platelet contamination

- The amount of lymphocytes is too high

- The amount of anti-HLA serum is too low

- Yellow colour of the HLA antisera

- Trays have been thawed and refrozen

- Reconstituted complement kept too long at room temperature before use

- Residual complement was frozen and thawed again

- Incubation times were too short

- Incubation temperature was too low

Causes of false positive reactions

- Cross reactions

- The amount of anti-HLA serum is too high

- Incubation times were too long

- Incubation temperature was too high

- Prior demage of lymphocytes (negative control is positive = „background“)

- Failure to add fixative

Rabbit Complement

Reconstitute lyophilized complement with 1 ml or 5 ml aqua dest. (according to package sizes). The reconstitution takes 10 - 15 minutes. Reconstituted complement must be stored cool (2...8°C) and used within 3 - 4 hours.

DO NOT FREEZE dissolved rabbit complement!

Performance characteristics

Please note the listing of test results HISTO TRAY ABC / AB / B27 RESERVOIR in order to receive data for diagnostic sensitivity and specificity (R-Value).

Literature

Bodmer, J. et al., 1997. Tissue Antigens 49:297-321

Warnings and Precautions

HISTO TRAY ABC /AB /B27 RESERVOIR are designed for in vitro diagnostic use only and should be applied by properly trained personnel, experienced in histocompatibility testing. Transplantation guidelines as well as EFI standard should be followed, in the particular case of doubtful typing results. Human source material used to produce these reagents has been tested and found negative for HBsAg and HIV and HCV antibodies. Nevertheless all used biological material like blood, sera and control sera should be handled as potentially infectious, because no test method can guarantee that material derived from biological sources are free from infectious agents. When handling biological material appropriate safety precautions are recommended (do not pipet by mouth; wear disposable gloves while handling biological material and performing the test; disinfect hands when finished the test).

Biological material should be inactivated before disposal (e.g. in an autoclave). Disposables should be autoclaved or incinerated after use. Spillage of potentially infectious materials should be removed immediatley with absorbent paper tissue and the contaminated areas swabbed with a suitable standard disinfectant or 70% alcohol. Material used to clean spills, including gloves, should be inactivated before disposal (e.g. in an autoclave).

Anti-HLA sera contain as preservative < 0.1% NaN3. A concentration of < 0.1% NaN3 is not considered to be a harmful concentration. Nevertheless avoid contact with the skin and mucous membranes. The copper and lead used in some plumbing systems can react with azides to form expolsive salts. The quantities of azide used in this reagent are small; nevertheless when disposing of azide-containing materials, they should be flushed away with a large volume of water.

Disposal of all specimen and test materials should be in accordance with state and local law.

For quenching solution, Formaldehyde solution and Acridinorange /Ethidiumbromide (AO/EB) please note the warnings and precautions of the manufacturers.

A yellow colouration of anti-HLA sera which still remains after thawing, may indicate a change of the pH value. Those plates should not be used for the test.

Material Safety Data Sheets available on request.

Do not use HISTO TRAY ABC / AB / B27 RESERVOIR beyond the indicated expiration date on the label.

Preservative: 	< 0.1% NaN3

Storage:	 - 20°C

Shelf life: 	until the expiration date indicated on the labels

Package:	according to information indicated on the labels

| # Explanation of symbols used on labbeling |  |
| --- | --- |
| IVD | For in vitro diagnostic use |
| LOT | Batch code |
| REF | Catalogue number |
|  | Storage temperature |
|  | Use by |
|  | Consult Instructions for use |

|  | BAG-BiologischeAnalysensystemGmbH Amtsgerichtsstr. 1 – 5, D-35423 Lich, Germany |
| --- | --- |
|  | Ordering Customer Service Phone: (+49) 64 04 / 9 25 450 (+49) 64 04 / 9 25 125 # Fax: (+49) 64 04 / 9 25 460 (+49) 64 04 / 9 25 421 # Internet: www.bag-germany.com |